## Slide 1: MEK-6400 series Technical seminar
Sensors adjustment

## Slide 2: Adjustment of the sensors
MEK-6400 series have 1 HGB measuring sensor, and 6 liquid detecting sensors.
  HGB measuring sensor
  WBC manometer sensor ( High, Low )
  RBC manometer sensor ( High, Low )
  Diluent and detergent detecting sensor
  Lysing reagent detecting sensor
This section is how to adjust these sensors.

## Slide 3: Adjustment on screen
Sensor voltage adjustment is available on the screen.
(*)except HGB sensor voltage
Notes: 試薬マネージメントが出来ます。

## Slide 4: HGB measuring sensor
HGB Measuring sensor

## Slide 5: Adjustment of HGB measuring sensor’s voltage
There is a HGB sensor with a WBC measurement bath in measurement unit MC-640V.
It is necessary to keep measurement reliability of HGB, the HGB sensor voltage should be adjusted around 3.0V in a state of HGB LED on.
The sensor voltage has to be less than 0.5V in a state of HGB LED off.
When the HGB sensor voltage is abnormal,
  “*” mark is displayed if the voltage is less than 0.5V when HGB LED is off.
  “!” mark is displayed if the voltage is more than 4.5V when HGB LED is on.
  “?” mark is displayed if the voltage is less than 1.5V when HGB LED is on.
In such a case, the reliability of HGB results becomes lower.

## Slide 6: Adjustment of HGB measuring sensor’s voltage
Adjustment method of the HGB sensor in the following.
Choose “OTHER” key from “MENU” screen. After that, choose “SENSOR MONITOR” function, so that you can see all sensors’ voltage on the screen.

## Slide 7: Adjustment of HGB measuring sensor’s voltage
Remove the side panel.
Check the WBC measurement bath is filled by diluent.
Turn the variable resistance of the HGB sensor voltage on UT-7201 ( HGB AMP BD, cf. illustration ) and adjust it so that the HGB sensor voltage becomes about 3.0V.
HGB sensor VR

## Slide 8: Adjustment of HGB measuring sensor’s voltage
Check HGB sensor voltage in a state of LED On is about 3.0V ( this figure is bad case )
Check HGB sensor voltage in a state of LED Off is less than 0.5V
After adjustment, replace the side panel.

## Slide 9: WBC/RBC Manometer sensor
WBC manometer (High)
WBC manometer (Low)
RBC manometer (Low)
RBC manometer (High)

## Slide 10: Adjustment of WBC/RBC Manometer sensors’ voltage
Two manometers for WBC and RBC absorb diluted sample of 360uL from apertures in measurement unit MC-640V.
Each manometer has two liquid surface detecting sensors, higher and lower sensor.
In the state that manometer is filled by diluent, the manometers sensors voltage should be lower than 1.5V.
In the state that manometer is empty, the manometers sensors voltage should be higher than 3.5V
A device judges to be or not to be of water at 2.5V. But if sensor voltages are between 2.25V - 2.75V, device cannot detect liquid surface precisely. To keep measurement reliability, device displays alarm such as “Manometer dirty“

## Slide 11: Adjustment of WBC/RBC Manometer sensors’ voltage
To adjust the manometer voltages, you have to change the operator to “Factory”
Choose “Setting” in “MENU”, and select “USERS“.
Choose "Factory" and push “CHANGE" key to change the operator.
You have to enter the password. Enter "4321“ as a password of "Factory".

## Slide 12: Adjustment of WBC/RBC Manometer sensors’ voltage
Adjustment method of the WBC/RBC Manometer sensors in the following.
Choose “OTHER” key from “MENU” screen. After that, choose “SERVICE” function.
Choose “MONITOR” key form “SERVICE“ screen, after that choose “Sensor monitor” function.

## Slide 13: Adjustment of WBC/RBC Manometer sensors’ voltage
Different point of two “Sensor monitor” screen, in normal menu and in service menu, is “INPUT CAL" key in service menu.
To adjust the manometer sensor voltages, push "INPUT CAL" key and change coefficient factors.

## Slide 14: Adjustment of WBC/RBC Manometer sensors’ voltage
Displayed contents are as follows.
WBC up:
Upper sensor coefficient of WBC manometer
WBC low:
Lower sensor coefficient of WBC manometer
RBC up:
Upper sensor coefficient of RBC manometer
RBC low:
Lower sensor coefficient of RBC manometer

## Slide 15: Adjustment of WBC/RBC Manometer sensors’ voltage
Move a cursor with arrow key and put “ENTER” key after the input with numerical value.
Push “OK” key to back to “Sensor monitor” screen, and check that the sensor voltage of manometers in filled by diluent is lower than 1.5V
And check that the sensor voltage is adjusted more than 3.5V in the state that there is no diluent in the manometer.
Repeat above operations till these voltages become reasonable values.

## Slide 16: Diluent /Detergent sensor, Lysing reagent sensor
Lysing
reagent
sensor
Diluent /
Detergent
sensor

## Slide 17: Adjustment of Diluent /Detergent, Lysing reagent sensor voltage
MEK-6400 has two reagent detection sensors which for diluent / detergent detection, and lysing reagent detection in “Inlet and outlet unit JQ-640V”.
Each liquid sensor voltage is less than 1.5V in a state to filled by reagent, and it is necessary to be adjusted than 3.5V in a empty state.
A device judges to be or not to be of water at 2.5V. But if sensor voltages are between 2.25V - 2.75V, device cannot detect liquid surface precisely.

## Slide 18: Adjustment of Diluent /Detergent, Lysing reagent sensor voltage
Displayed contents are as follows.
Diluent:
Diluent and Detergent detecting sensor coefficient.
Lysing:
Lysing reagent detecting sensor coefficient of WBC manometer

## Slide 19: Adjustment of Diluent /Detergent, Lysing reagent sensor voltage
Move a cursor with arrow key and put “ENTER” key after the input with numerical value.
Push “OK” key to back to “Sensor monitor” screen, and check that the sensor voltage in filled by reagent is lower than 1.5V
And check that the sensor voltage is adjusted more than 3.5V in the state that there is no diluent.
Repeat above operations till these voltages become reasonable values.